## Slide 1: Validation Services

## Slide 2: Andreas Kößler
In Ellab since March 2019
Positions:
Area Sales Manager
Head of Sales & Marketing DACH
Country Manager CH
Validation Engineer
Brand Manager Director
Validation Solutions

## Slide 3: Intro to Validation Services Part 1
What is it that our customer needs
Ask! Ask! Ask!
”Do not anticipate the answer”
Customer needs /Applications

## Slide 4: Overview of Units and Environment Under Test
Applications
  Controlled Temperature Units (CTUs)
    Liquid Nitrogen (LN2) Freezers
      Vapor and Liquid Phase
    -150°C Mechanical Freezers
    -90°C Ultra-Low to -20°C Freezers
    Refrigerators (2°C to 8°C)
  Environmental chambers
    Incubators: Temperature, Humidity and CO2, Shakers
    Stability chambers: Temperature, Humidity and Photostability
  Controlled Temperature Environments (CTEs)
    Stability storage rooms, Walk-In Cold Rooms, Walk-In Freezers
  Warehouse Mapping
  Autoclaves (EN285) and Sterilizers
  Lyophilizers / Freeze dryers
  Depyrogenation Ovens & Tunnels
  ETO Sterilization
  Vaporized H2O2 processes
  Steam-In-Place (SIP) processes
  Centrifuges, Water baths and Washers
  + much more
Qualification Services of units and environments

## Slide 5: Types of Services Commonly Required
Consulting, short-term, and long-term staffing
New site and site expansion support
On-Site Field Calibrations
FAT, SAT, Commissioning
Qualifications: CTUs, CTEs, Warehouses, Storage Areas
Validations: Autoclaves, Ovens (Various), Controlled Rate Freezers, Glass Washers, Lyophilizers
Additional Services: Steam Quality Testing, Monitoring Support
Customer needs

## Slide 6: Types of On-site Thermal Validation / Mapping Services
Customer needs
  Design Qualification (DQ)
  Engineering Studies
  Installation Qualification (IQ)
  Operational Qualification (OQ)
  Performance Qualification (PQ) or Continuous Qualification (CQ)
  Instrument / equipment calibration
  Power off / Power failure tests
  Door Open tests
  Temperature Distribution tests
  Temperature Penetration tests
  Alarm testing
  Loaded and Empty Mappings
  Shelf mapping, vacuum testing, Bowie Dick testing, steam quality testing, vacuum leak test, conductivity tests, riboflavin removal testing, controlled rate mapping
  And several more…

## Slide 7: Why choose Ellab to Perform On-site Services?
Customer needs
  Ellab is committed to providing world-class services to meet all our customer’s qualification, validation and calibration needs
  We offer a full range of field services for GMP/GLP facilities in the Life Science and Food Manufacturing Industries
  Our goal is to build long-term relationships with customers
  We differ from competitors through our relationships with customers, dedicated and hardworking staff, expertise and efficiency in the services we perform
  Ellab is committed to delivering high quality, which when combined with Ellab's equipment, leads to excellent qualification and validation work

## Slide 8: Typical Challenges Ellab On-site Services can Solve
Solution
Examples of Challenges | Ellab’s On-site Services
Projects may come up quickly | We provide an appropriate turnaround time on proposal and manage expectations Have multiple locations with responsive engineers
Lack of specialized and trained staff | We focus on mid to sr. level engineers that are capable of qualifying equipment efficiently with a high level of quality Newer employees work in tandem with more experienced engineers till ready to work solo
Projects can vary in terms of size and turnaround times | Industry leading equipment for quality & efficiency Quantity can be optimized for each project Easy to transport Lower shipping costs Can fly with equipment on smaller projects
Cumbersome protocols Slow turnaround time of reports Lack of quality, detail, clarity on reports. Lack of GDP (Good Documentation Practices) | Streamlined protocols that are customized utilizing autofill areas for customer & unit information Easy integration of ValSuite reports

## Slide 9: Services we offer
Applications
Staffing of Service Engineers
Calibration Services
Monitoring Services
  Thermal Qualifications/Validation Projects
    Ellab offers our highly skilled engineers to perform on-site thermal qualification and other associated validation tasks
    Available turn-key validation projects performed in a high quality / efficient manner including protocol generation, field execution, and summary report packages
    Ability to support scopes of work for any size project
  GMP Staffing & Projects
    Our engineers can also perform various other GMP related tasks
    Long term contracts available – Hourly, Weekly, Monthly, Annual contacts available
  Skilled Calibration Engineers with High Quality Equipment
    Calibrations outside a conditioned metrology laboratory require a high level of field calibration knowledge and experience
    With Ellab you have a competent partner for these comprehensive calibration tasks of your high-grade instruments and measuring systems
  On-site Calibration Services
    Control sensors on autoclaves, Pasteurization plants, UHT facilities and Cold rooms
    Temperature Calibration range from -200 to +400°C plus
  IQ / OQ of new monitoring systems
    An essential part of acquiring a new monitoring system is have it qualified
    Ellab offers both the monitoring systems and the required IQ / OQ services
  IQ / OQ of additional monitoring units
  Re-calibration of monitoring units

## Slide 11: Intro to Validation Services Part 2
Our offering & How to quote

## Slide 12: Foundational concepts:
“The better you know yourself, the better your relationship with the rest of the world”
-Toni Collette
“The most important knowledge is understanding what you can’t do.”
-Bill Buford

## Slide 13: Types of Services Offered
Consulting
On-Site Field Calibrations (Various depending on standards and staff available)
Commissioning, FAT, SAT
Qualifications: CTUs, CTEs, Warehouses, Rooms
Validations: Autoclaves, Ovens (Various), Controlled Rate Freezers, Glass Washers, Lyophilizers
Additional Services: Steam Quality Testing, Monitoring Support
Short-term, and long-term staffing
Hourly projects and fixed price projects
Our offering will vary from region to region

## Slide 14: Core Concept – Determining and Valuing Time
Time – Understanding the SOW to derive the expected quantity of hours, days, weeks, or months it will take to complete a particular SOW
    Could be for one staff member, or for many.
A client says “Please calibrate and qualify my refrigerator”
    Does that mean 2 to 3 days of work, or over a week’s worth of work? If we estimate incorrectly that could be either devastating to the company if we win and/or lose the project.
What are the expectations around efficiency, and how does that effect your pricing methodology?
    Can you only qualify one unit at a time, or 4?
Time can vary even with similar sounding SOW (Scope of Work)

## Slide 15: Core Concept – Materials
What needs to be provided by Ellab to complete the project at hand.
    Calibration equipment
    Ellab mapping equipment
      Data Loggers, E-Val Pro, Sky, TVP, or our choice
      Equipment fees built into fixed/hourly rate, or separate charge?
      What if the prospect wants you to use their equipment?
    Steam quality test kits
    Other materials
      Consumables: Binders, paper…etc.
      Calibration standards, buffers…etc.
      Shipping materials
      TIP: Price per measurement point/day. E-Val Pro = 3Points
Material can, but should never be a showstopper

## Slide 16: To hourly or not to hourly?

## Slide 17: Hourly Price Projects (1/2)
Staffing & Hourly projects with/without equipment (Small/Medium/Large)
Complex projects
Some calibration projects are better as hourly (Rare)
Validations
    Autoclaves, SIP, Ovens, Glasswashers…etc.
Utilizing client’s quality system
    Protocol and report system = Generation of documents in their system
Mapping equipment
    Part of hourly rate, separate fee/rental, or no equipment required
Quantity of units under test available for simultaneously mapping may be limited
Large potential for delays
    Too many variables outside of our control
    Working with or through other contracting firms
Always safe > $$$ 

## Slide 18: Hourly Price Projects (2/2)
Hourly projects – Autoclaves, Ovens, SIP…etc.
    Easiest – They tell you.
    Likely – You need to figure out. Work out with your validation team to determine a game plan and get agreement with the prospect.
    The larger and more complex projects are more difficult to gauge
    Strategy: Cannot come in too low, but also cannot come high as those are often
    automatically disqualified
Staffing roles
    They need to tell you the duration because your staff will essentially become part of their staff for a fixed time period.
    Is typically straight forward and most likely it takes longer as expected.
How do you figure out how many hours?

## Slide 19: Example of hourly quote

## Slide 20: Fixed Price Projects
Requires a well-defined SOW
Are there:
  Training requirements
  Calibrations to be executed
  Protocol to be prepared filled in
Qualifications / Validations
  CTUs/CTEs, warehouses, Steam Quality Testing
    Defined quantity of multiple units qualified simultaneously
    Mapping equipment
Summary Report Package
TIP: Put disclaimer in quote just in case things are not like promised 
If all is well planned and predicted-> $$$ 

## Slide 21: Protocol Options
Ellab Protocol
    Does it require any customization?
    Build pricing into per unit, or separate line item?
Client’s protocol
    Could be simple, lean, extensive, or incredibly complex
Manufacturer’s protocol
    Often not the most efficient protocol to execute
Field Test / Commissioning document
    Simple 1 page document with basic tests for FAT, SAT, commissioning level testing.
Protocols are hidden money loosing monsters 

## Slide 22: Determining the Minimum Quantity of Mapping Sensors: Reach-Ins and Walk-ins
Minimum ambient sensors based on ISPE guide
Chamber volume: X ≤ 70 cubic feet (2 cubic meters)
  9 sensors
Chamber Volume: 70 < X ≤ 700 cubic feet (20 cubic meters)
  15 sensors
*Actual number of mapping sensors may vary and often exceed the minimums based on the ISPE guide.
Common quantities of ambient sensors seen in reach-in sensors
  Temperature only: 12
  Temperature / RH: 12 Temperature, 3 RH
  Temperature / RH / CO2: 12 Temperature, 3 RH, 1 CO2
Common quantities seen in walk-in quantities (will vary depending on size)
  Temperature only: 16+
  Temperature / RH: 16+ Temperature, 4+ RH
ISPE Good Practice Guide: Controlled Temperature Chamber Mapping and Monitoring

## Slide 23
How long does it take on-site? Single Set Point, 24-hour mappings
IQ
Controller OQ (Alarm Tests, Power Interruption)
OQ Mapping (24-hour mapping, Open door Test)
Calibration
3ish days
IQ
Controller OQ (Alarm Tests, Power Interruption)
OQ Mapping (24-hour mapping, Open door Test)
Calibration
4ish days
PQ Mapping (24-hour ‘loaded mapping)
IQ
Controller OQ (Alarm Tests, Power Interruption)
OQ Mapping (24-hour mapping, Open Door Test, Power Loss Mapping)
Calibration
5 days
PQ Mapping (24-hour ‘loaded mapping, Open Door Test, Power Loss Mapping)
Standard Cal / IOQ
Standard Cal / IOPQ
Fully Loaded Cal / IOPQ
Requal Mapping (24-hour ‘as-is’)
Calibration
2ish days
Standard Cal / Requal

## Slide 24: How long does it take on-site? 48-hour mappings
IQ
Controller OQ (Alarm Tests, Power Interruption)
OQ Mapping (48-hour mapping, Open door Test)
Calibration
4ish days
IQ
Controller OQ (Alarm Tests, Power Interruption)
OQ Mapping (48-hour mapping, Open door Test)
Calibration
6~7ish days
PQ Mapping (48-hour ‘loaded mapping)
IQ
Controller OQ (Alarm Tests, Power Interruption)
OQ Mapping (48-hour mapping, Open Door Test, Power Loss Mapping)
Calibration
7~8ish days
PQ Mapping (48-hour ‘loaded mapping, Open Door Test, Power Loss Mapping)
Standard Cal / IOQ
Standard Cal / IOPQ
Fully Loaded Cal / IOPQ
Either 2 set points, or 2 conditioners
Requal Mapping(48-hour ‘as-is’ mapping)
Calibration
3ish days
Standard Cal / Requal

## Slide 25: Typical Project Workflow
A good planning always helps
Request
Verify SOW
Provide Recommendation
Order
Introduction to customer
Discuss schedule & Send out information sheet
Execution
Schedule project
Execute work
Protocol Approval
Send protocol for review and approval
Project Completion
Complete and ship Summary Report Package
Verify Mapping Equipment

## Slide 26: Get our Hands dirty
Customer is asking the following:
I have got 11 ULT - Freezers with two setpoints at -86°C and -65°C.
I need them to be initially qualified asap. Including all.
Could you send me a quote please?
Oh if you are on it also quote the yearly requal (24h mapping only) for all of them.
As we are in the same town – You are my guys!
You quote

## Slide 28: Intro to Validation Services Part 3
Sterisense
SteriSense®

## Slide 29: Key Advantages of SteriSense
The Ultimate Electronic Bowie Dick Test
Accurate Electronic Bowie Dick Test
Reliable and cost efficient
Very compact size
User-friendly and intuitive software
Eliminates guesswork and provides and objective result
Replaceable PCD to run several cycles in a sequence
Verification of critical sterilization parameters beyond traditional methods
Up to 1,000 test cycles between each calibration

## Slide 30: Why use SteriSense® Instead of Traditional Indicators?
The Ultimate Electronic Bowie Dick Test
SteriSense provides a much more in-depth verification of critical sterilization parameters, eliminating any guess work.
The Software also automatically generates and stores printable reports, ensuring electronic documentation of the process.
In summary, by using SteriSense you gain the following advantages compared to chemical or biological indicators:
  Economically efficient with high test volumes
  Improved safety - objective results that eliminate the risk of false or grey zone readings
  Additional critical information - checks performed according to ISO 17665
  Easy to store, retrieve and compare the data from a database - convenient and safe method that eliminates the need for guessing
  Time stamps included to increase reliability

## Slide 31: Why use SteriSense® Instead of Other Electronic Bowie Dick Test Devices?
The Ultimate Electronic Bowie Dick Test
SteriSense is by far the most compact Electronic Bowie Dick Test on the market.
The Process Challenge Device (PCD) has the unique feature of being replaceable, which allows the user to run several test cycles in a sequence with very limited need for cooling.
This reduces the amount of time and investments spent on unnecessary backup devices.
The replaceable PCD feature is patent pending.
In summary, by using the BD sensor you gain the following advantages compared other electronic Bowie Dick Tests
  Ultra-compact size
  Highly user-friendly software
  Advanced sterilization analysis
  Environmentally friendly process
  Up to 1,000 test cycles between each calibration

## Slide 32: Why use SteriSense® Instead of Other Electronic Bowie Dick Test Devices?
The Ultimate Electronic Bowie Dick Test
The PCD has been specially designed to reflect the reference method originally developed by Dr. J. Bowie and J. Dick in the 1960’s
The function of the PCD is to “challenge” the steam penetration of a steam sterilizer in accordance with ISO 11140-4,134˚C & 121°C
The two temperature sensors measures the temperature inside the PCD and the ambient temperature in the sterilization chamber, of which the values have to follow each other
The pressure sensor measures the ambient pressure value within the chamber, which is used for detecting vacuum pulses, calculating the theoretical ambient temperature and detecting super heated steam
It is also possible to check for the presence of non-condensable gases (NCG) by analysing the steam’s capability to heat the steel core inside the PCD
The checks performed in the Master Session is referring to ISO 17665