| Autoclave & Steam Quality Questionnaire Version V1.0.1 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Territory Manager: |  | Your Name | Date Submitted: | DD-MMM-YYYY |
| Customer and project location: Company Name Street City, State Zip Code Attention: Contact Name Phone Number (Office): Phone Number Phone Number (Cell): Phone Number E-mail Address: E-mail Address Project Location: If different than contact location |  |  |  |  |
| ValSuite validation requirements: |  |  |  |  |
| Do they require the ValSuite we are using on-site to be validated? If so, do they require a copy of a validation binder? |  |  | If no validation, then no worries. If yes to validation, but no binder needed that is OK. If need binder, additional fee will apply. |  |
| Quantity of Autoclaves: |  |  |  |  |
| Quantity of Autoclaves: | (Based on complexity, may need additional questionnaires to keep units separate) |  |  |  |
| Unit Calibrations: |  |  |  |  |
| Are we calibrating their Autoclave? |  | Yes/No: | If yes, is a single point (point of use) calibration acceptable, or do they require a 3 point calibration? | Single Point or Three Point? As-Found & As-Left? As-Left only? |
| Protocols: (Pick One or the other) |  |  |  |  |
| Who’s protocol are we executing? |  |  | Option 1: Ellab Option 2: Client’s protocol – Price reduction if using their protocol |  |
| Final Reports: (Pick One or the other) |  |  |  |  |
| Who will be generating the final report? |  |  | Option 1: Ellab. Option 2: Client - Price reduction if client does final report |  |
| Review of Final Report: (Pick One or the other) |  |  |  |  |
| Ellab review: |  |  | Yes/No: Included in standard cost (Client signs off the Review By on the protocol) |  |
| Ellab Quality Department Review: |  |  | Yes/ No: Additional cost & turnaround time (Estimated $95~$125 + couple more days |  |
| Qualifications: (Make the appropriate selections below |  |  |  |  |
| Level of qualification: | Commissioning CGMP/ISO17665 EU (EN 285) |  |  |  |
| Vacuum or Gravity only unit: | (Vacuum & Gravity) or Gravity only |  |  |  |
| Any development work necessary? | Yes/No (Development work happens for new autoclaves, or new cycles being added to an existing autoclave) |  |  |  |
| Cycle Development studies | If Yes, is formal protocol/documentation required? Yes/No |  |  |  |
| Developing cycle: | Number or estimated runs, dependent typically on the number of cycles. |  |  |  |
| Installation Qualification: | Yes/No: |  |  |  |
| Operational/Re-Qualification: | Yes/No: Add more cycles (rows) if necessary. |  |  |  |
| Cycle Type / Runs: | - Gravity 121.1C / Run(s) - Pre-Vacuum and/or Post Vacuum 121.1C / Run(s) - Liquids 121.1C / Run(s) - Flash 132C to 135C / Run(s) |  |  |  |
| Performance/Re-Qualification: | Yes/No: (Testing with BIs) |  |  |  |
| Flexible Load Approach: | Add more cycles (rows) if necessary. One row for each type of load. |  |  |  |
| Item Load Mapping Runs: | For Flexible Load Approach. Number of items needing to be mapped. Rule: Approximately 10 to 12 items per run |  |  |  |
| Cycle Type / Load Type / Duration / Runs: | - Gravity 121.1C / (Glassware, unwrapped goods, waste, utensils, red-bags) / (Fractional/Half), Full, / # Minimum Run(s) and # Maximum Run(s) - Pre-Vacuum and/or Post Vacuum 121.1C / (Glassware, Wrapped goods, packs, porous materials, red-bags) / (Fractional/Half), Full, / # Minimum Run(s) and # Maximum Run(s) - Liquids 121.1C / (Growth media, LB Broth, Water…etc) / (Fractional/Half), Full, / # Minimum Run(s) and # Maximum Run(s) - Flash 132C to 135C / (Unwrapped goods…for shorter periods) / (Fractional/Half), Full, / # Minimum Run(s) and # Maximum Run(s) |  |  |  |
| Duration of run: | Fractional/Partial: A the duration that should inactive the BIs, and provides rationale for overkill strategy (Typically dry goods) or Full: Run the full length of the product specific programmed cycle |  |  |  |
| Fixed Load Approach: | Add more cycles (rows) if necessary |  |  |  |
| Cycle Type / Load Type / Duration / Runs: | - Gravity 121.1C / (Glassware, unwrapped goods, waste, utensils, red-bags) / (Fractional/Half), Full, / # Run(s) Fixed - Pre-Vacuum and/or Post Vacuum 121.1C / (Glassware, Wrapped goods, packs, porous materials, red-bags) / (Fractional/Half), Full, / # Run(s) Fixed - Liquids 121.1C / (Growth media, LB Broth, Water…etc) / (Fractional/Half), Full, / # Run(s) Fixed - Flash 132C to 135C / (Unwrapped goods…for shorter periods) / (Fractional/Half), Full, / # Run(s) Fixed |  |  |  |
| Additional Testing: |  |  |  |  |
| Bowie Dick Tests with Warmup: | Yes/No: Vacuum air removal test ensuring enough air is removed. If yes, how frequently? Beginning of day? Only after TC incursion? |  |  |  |
| Vacuum Leak Tests: | Yes/No: Air-tight integrity (Typically after TC incursion). If yes, how frequently? Beginning of day? Only after TC incursion? |  |  |  |
| Mapping Equipment: |  |  |  |  |
| Are they OK with us using Data Loggers? |  |  | Yes/No |  |
| Do they require the use of thermocouple based data loggers? |  |  | Yes/No |  |
| Timeline: |  |  |  |  |
| Project Start Date On-Site: When do they want it to start, reasonably. Does not have to be an exact date (EG, Early September 2016) |  |  |  |  |
| Project plan and/or milestones: If we are going to have all the autoclaves available simultaneously, or mixed with production? If a larger project then there may be specific milestones they want completed by certain dates. |  |  |  |  |
| Project Completion Date On-Site: When do they want everything wrapped up by on-site, and to receive completed protocols by? We will often wrap up the work on-site and then finish the documentation at our office instead of adding to the travel expenses. |  |  |  |  |
| Any site specific training required?: (Client specific training in order for our team to perform work at their site) |  |  |  |  |
| If no, then just say none here. If yes, describe. Couple hours, half day, whole day, multiple days, weeks… |  |  |  |  |
| Other: |  |  |  |  |
| Anything else we should know? |  |  | Yes/No and if yes, then what is it? |  |
| Competitors? |  |  |  |  |
| Previous pricing strategy (Fixed/Hourly)? |  |  |  |  |
| Steam Quality Testing: Typically for EN 285 |  |  |  |  |
| How many drops: Unit 1 |  |  |  |  |
| How many drops: Unit 2 |  |  |  |  |
| Types of fittings: |  |  |  |  |